## Slide 1
Announcements
HW: Due Tuesday Oct 27th.
Start looking at an article to do your write-up on.
Best article is from Nature, Science, Cell.
(though can choose from any scientific journal)
Want an article that has a significant impact on field.
Thursday, Oct 29th, turn in the article, a secondary article that is a review of the field at which your article is about.
Turn in a paragraph, summarizing it.
The article’s general point, and why you’ve chosen it.
(why it’s important, interesting,
Send your write-up and PDF of article to me by email.
(It won’t be graded but will make sure it’s appropriate.)

## Slide 2
Finish up Optical Traps
Move on to:
Seeing things with Microscopy
…Mostly fluorescence
Why can you see single particles
and get nanometer resolution
even with visible light

## Slide 3
Basic Optical Trap set-up
http://en.wikipedia.org/wiki/Optical_tweezers
Optical Trap/ Tweezer

## Slide 4
Requirements for a quantitative optical trap:
1) Manipulation – intense light (laser), large gradient (high NA objective), moveable stage (piezo stage) or trap (piezo mirror, AOD, …) [AcoustOptic Device- moveable laser pointer]
2) Measurement – collection and detection optics (Back Focal Plane interferometry)
3) Calibration – convert raw data into forces (pN), displacements (nm): Brownian Test force.

## Slide 5
Measuring the position of a trapped bead
Want a Position Sensitive Detection to measure

## Slide 6
How you get parallel light?
Put object at the focal length of lens. Then image is at infinity.
Recall Len’s Maker’s Equation

## Slide 7
2) Measurement: Back Focal Plane imaged onto detector
Trap laser
BFP
specimen
PSD
Relay lens
Conjugate image planes
∑
(Image at point P will
get imaged to point P’)
Notes: Typically 1000 nm deflection for 1V of signal
10 uV noise is 0.1A !!!

## Slide 8
N
P
N
P
In1
In2
Out1
Out2
Position sensitive detector (PSD)
Plate resistors separated by reverse-biased PIN photodiode
Opposite electrodes at same potential
– no conduction with no light

## Slide 9
N
P
N
P
In1
In2
Out1
Out2
ΔX ~ (In1-In2) / (In1 + In2)
ΔY ~ (Out1-Out2) /(Out1+Out2)
POSITION
SIGNAL
Multiple rays add their currents linearly to the electrodes, where each ray’s power adds Wi current to the total sum.
Linear signal with position.
Can just read off signal, get position

## Slide 10
Laser
Beam expander
Objective
Condenser
Photodetector

## Slide 11
The noise in position using equipartition theorem
[For a typical value of stiffness (k) = 0.1 pN/nm.]
<x2>1/2 = (kBT/k)1/2 = (4.14/0.1)1/2 = (41.4)1/2 ~ 6.4 nm
What is noise in measurement?.
6.4 nm is a pretty large number.
[ Kinesin moves every 8.3 nm; 1 base-pair = 3.4 Å ]
How to decrease noise?
Reduce bandwidth. (Get to soon!)
Also: Operate at high force less noise due to finite Temp.
Also be clever about how to differentiate noise from signal.
Hint: Equipartion Theorem calculates for noise at all frequencies (infinite bandwidth).
Notes: If you look at the expression for the power spectrum, amplitude at low frequency (basically the part which is flat in frequency) is 4*kT*gamma/k^2. (The alpha term is if you measure the power spectrum in V (volts): assume everything is converted into position (in nm), then you don’t need the alpha^2.)
To get the position noise, you need to integrate the power spectrum over frequency. If the bandwidth BW is smaller than the corner frequency, then it’s a simple integration, because the power spectrum is constant, with amplitude 4*kT*gamma/k^2.
So, the noise over bandwidth BW is 4*kT*gamma*BW/k^2.
To get the rms noise in the bandwidth, you take the square root: sqrt(4*kT*gamma*BW)/k
For a BW = 100 Hz, gamma = 1e-6, stiffness k = 0.1 you have: sqrt(4*4.14*1e6*100)/0.1
(all in units of pN, nm)

## Slide 12
Also, know how to cut out noise.
Take out frequencies where signal isn’t
Solve Langevin’s Equation,
Get power spectrum of bead in an optical trap.
γ = 6πηr (for sphere)
Noise is not distributed evenly across all frequencies in an optical trap. Most noise at low f.
= trap stiffness
Bandwith = infinite:
limit to ~6 nm.
If use BW = 100 Hz
~ 0.4 nm = 4 Angstrom!!
Notes: Noise = integrate power spectrum over frequency.
If BW < fc then it’s simple integration because power spectrum is constant, with amplitude = 4kBTg/k2

## Slide 13
3.4 kb DNA
F ~ 20 pN
f = 100Hz, 10Hz
1bp = 3.4Å
1
2
3
4
5
6
7
8
9
1
2
3
4
5
6
7
8
9
UIUC - 02/11/08
Basepair Resolution—Yann Chemla @ UIUC
PNAS
Take difference and sum.
Diff: no shaking of floor!
Limit BW,
Eliminate noise

## Slide 15
Calibrate w Brownian motion as test force
Drag force
γ = 6πηr
Fluctuating
Brownian force
Trap force
<F(t)> = 0
<F(t)F(t’)> = 2kBTγδ (t-t’)
kBT= 4.14pN-nm
Langevin equation:
Inertia
term
(ma)
≈0
Inertia term for mm-sized objects is always small
(…for bacteria)
Can solve equation
Notes: kT = 1.38e-23*300 = 4.14e-21 = 4.14pN nm
Sqrt(<x^2>) = sqrt(kT/k) = sqrt(4.14/0.1) = sqrt(41.4) ~ 6 nm

## Slide 16
Class evaluation
What was the most interesting thing you learned in class today?
2. What are you confused about?
3. Related to today’s subject, what would you like to know more about?
4. Any helpful comments.
Answer, and turn in at the end of class.

## Slide 17
Calibrate w Brownian motion as test force
Drag force
γ = 6πηr
Fluctuating
Brownian force
Trap force
<F(t)> = 0
<F(t)F(t’)> = 2kBTγδ (t-t’)
kBT= 4.14pN-nm
Langevin equation:
Inertia
term
(ma)
≈0
Inertia term for mm-sized objects is always small
(…for bacteria)
Exponential autocorrelation function
Notes: kT = 1.38e-23*300 = 4.14e-21 = 4.14pN nm
Sqrt(<x^2>) = sqrt(kT/k) = sqrt(4.14/0.1) = sqrt(41.4) ~ 6 nm